## Data-Schema-Creep-v2.0
CATEGORY I | CATEGORY II | CATEGORY III | CATEGORY IV | ENTRY | ENTRY - ADDITIONAL INFORMATION | SYMBOL | UNIT | DATA TYPE | EXEMPLARY ANSWER OR OPTIONS SEPARATED BY SLASH IN CASE OF DATA TYPE DROP-DOWN LIST | REQUIREMENT | COMMENTS
Metadata | Test info | Test job details |  | Date of test start |  |  |  | String | 2023-04-25 | Mandatory
Metadata | Test info | Test job details |  | Date of test end |  |  |  | String | 2023-05-15 12:00 | Mandatory
Metadata | Test info | Test job details |  | Test ID |  |  |  | String | Vh5205-C-99 | Mandatory
Metadata | Test info | Test job details |  | Project |  |  |  | String | NFDI-MatWerk | Optional
Metadata | Test info | Test parameters |  | Test standard applied | Was the test performed according to a test standard? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory
Metadata | Test info | Test parameters |  | Test standard |  |  |  | Drop-down list | ISO 204 / ASTM E139 / Other (Specify) | Mandatory | Applicable if test performed according to a test standard (row 6)
Metadata | Test info | Test parameters |  | Specified temperature |  | T | ℃ | Numeric | 980 | Mandatory
Metadata | Test info | Test parameters |  | Type of loading |  |  |  | Drop-down list | Tension / Compression / Bending | Mandatory
Metadata | Test info | Test parameters |  | Load control type |  |  |  | Drop-down list | Constant force / Constant stress | Mandatory
Metadata | Test info | Test parameters |  | Initial stress |  | Ro | MPa | Numeric | 170 | Mandatory | Applicable if test performed under constant force (row 10)
Metadata | Test info | Test parameters |  | Test type |  |  |  | Drop-down list | Uninterrupted creep tests with continuous monitoring of extension / Interrupted creep tests with periodic measurement of elongation / Stress rupture tests where normally only the time to fracture is measured | Mandatory
Metadata | Test info | Test parameters |  | End of test criterium |  |  |  | Drop-down list | Test piece break / Time limit / Extension limit / Other (Specify) | Mandatory
Metadata | Test info | Test parameters |  | Time Limit |  |  |  | String | 100 h | Mandatory
Metadata | Test info | Test parameters |  | Extension Limit |  |  |  | String | 1% creep extension | Mandatory
Metadata | Test info | Test parameters |  | End of test | Describe the end of test if none of the preset criteria (time or extension limit) could be met |  |  | String |  | Optional
Metadata | Test info | Test parameters |  | Interruption course |  |  |  | Drop-down list | Unloading before cooling / Unloading after cooling | Mandatory | Applicable if test was not performed until test piece break (Row 13)
Metadata | Test info | Test parameters |  | Test force |  |  | kN | Numeric | 5.1 | Optional
Metadata | Test info | Test parameters |  | Preload | Part of the test force |  | kN | Numeric | 0.2 | Mandatory
Metadata | Test info | Test parameters |  | Additional information |  |  |  | String |  | Optional
Metadata | Test info | Related research outcome |  | Any related articles |  |  |  | Drop-down list | Yes / No | Optional
Metadata | Test info | Related research outcome |  | Related article | Multiple input entries for instance for MSE research article DOI, Zenodo DOI, FDO |  |  | String, Array | https://doi.org/10.1080/09603409.2016.1186414 | Optional
Metadata | Material history and condition |  |  | Material Identifier | E.g., NIMONIC 75, 2.4630, CMSX-6, CMSX-4, ERBO1, … |  |  | String | NIMONIC 75 | Mandatory
Metadata | Material history and condition | As-manufactured material |  | Phase transformation during test |  |  |  | Drop-down list | Yes / No | Optional
Metadata | Material history and condition | As-manufactured material |  | Possible phase transformation | Is any phase transformation expected due to temperature during the creep test?. Please provide any supporting material if possible. Answers could be, e.g., a link to a TTT-curve or a DOI of an article |  |  | String | Link to a TTT-curve | Optional
Metadata | Material history and condition | As-manufactured material |  | Form of as-manufactured material | E.g., Cast, Ingot, Extrusion rod, … |  |  | String | Cast | Mandatory
Metadata | Material history and condition | As-manufactured material |  | Geometry/size as-manufactured material | Please provide a description or/and any supporting material, e.g., link to image or technical drawing, if possible. |  |  | String | Plate (140 mm x 100 mm x 20 mm) | Mandatory
Metadata | Material history and condition | As-manufactured material |  | Manufacturing process description as-manufactured material | E.g., Cast / Melting, Casting, and Remelting / Induction melting in air, casting into a circular ingot and then electroslag remelting |  |  | String | Melting, Casting, and Remelting | Optional
Metadata | Material history and condition | As-manufactured material |  | Casting temperature |  |  | °C | Numeric, Array |  | Optional
Metadata | Material history and condition | As-manufactured material |  | Casting speed |  |  |  | Numeric, Array |  | Optional
Metadata | Material history and condition | As-manufactured material |  | Solidification | Single/Polycrystal solidified? |  |  | Drop-down list | Single crystal / Polycrystal | Mandatory
Metadata | Material history and condition | As-manufactured material |  | Condition |  |  |  | Drop-down list | Heat treated / As manufactured | Mandatory
Metadata | Material history and condition | As-tested material |  | Supplier |  |  |  | String | Doncasters Precision Castings-Bochum GmbH | Mandatory
Metadata | Material history and condition | As-tested material |  | Geometry/size as-tested material | The test piece is manufactured from the as-tested material. Please add a description or a link to Image or technical drawing. The as-tested material is the material to be tested. The as-tested material can be a component. |  |  | String | Answers could be, e.g., Plate of 30 mm, or rectangular section taken from pipe wall | Mandatory
Metadata | Material history and condition | As-tested material |  | Manufacturing process description as-tested material |  |  |  | String | Cut out from the cast plate. No further heat treatment | Mandatory
Metadata | Material history and condition | As-tested material |  | Supply Date |  |  |  | String |  | Mandatory
Metadata | Material history and condition | As-tested material |  | Order number |  |  |  | String | ERBO 1A Item 1 | Optional
Metadata | Material history and condition | Heat treatment |  | Heat treatment - State |  |  |  | Drop-down list | None / Annealed / Hardened / Other (Specify) | Mandatory | Applies if row 32 = Heat treated
Metadata | Material history and condition | Heat treatment |  | Multistage annealing? |  |  |  | Drop-down list | Yes / No / Not applicable | Optional
Metadata | Material history and condition | Heat treatment |  | Multistage ageing? |  |  |  | Drop-down list | Yes / No / Not applicable | Optional
Metadata | Material history and condition | Heat treatment |  | Heat treatment - Annealing - Description |  |  |  | String | Heat-up at 20 °C per min auf 1290 °C, Hold at 1290 °C 1h, Heat-up at 1°C per min to 1300 °C - 1315 °C, Hold at 1300 °C - 1315 °C for 6h. Cool down at 150 °C - 400 °C per min to ≤ 800 °C. Further cooling as required. | Mandatory | Applies if row 38 = Annealed
Metadata | Material history and condition | Heat treatment |  | Heat treatment - Ageing - Description |  |  |  | String | Under inert gas at (1140 ± 10) °C 4 h / cooling down to 800 °C under inert gas, below 800 °C under inert gas or in air then: Under inert gas at (870 ± 10) °C 16 h cooling down to 800 °C under inert gas; below 800 °C to room temperature under inert gas or in air. | Mandatory | Applies if row 38 = Hardened
Metadata | Material history and condition | Heat treatment |  | Heat treatment - Protocol | Link to file, preferably with machine-readable (meta)dataeta)data |  |  | String | Link to protocol file | Optional
Metadata | Material history and condition | Microstructure |  | Microstructure feature |  |  |  | Drop-down list | Matrix / Phase / Grain Boundary / Dendrite / Precipitate / Inclusion / Grain / Segregation / Other (Specify) | Mandatory
Metadata | Material history and condition | Microstructure |  | Characterization method | E.g., STEM, HRTEM, TEM-EDX, AFM, SEM, SEM-EBSD, …. |  |  | String | SEM-EBSD | Mandatory
Metadata | Material history and condition | Microstructure |  | Measured condition | E.g., as-manufactured, heat-treated, before testing, after testing |  |  | String |  | Mandatory
Metadata | Material history and condition | Microstructure |  | Measuring position | Specify position(s) for bulk (middle/top/bottom/surface) and/or microstructure feature (matrix, interdendtitic region, phase,...) |  |  | String |  | Mandatory
Metadata | Material history and condition | Microstructure |  | Microstructure feature - Information | Include all relevant characterization results |  |  | String, Array |  | Optional
Metadata | Material history and condition | Microstructure |  | Microstructure Image | Link to File, e.g., an optical micrograph, preferably with machine-readable (meta)data |  |  | String | Link(s) to image(s) (can be from different characterization techniques) | Optional | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Microstructure |  | Microstructure report | Link to file, preferably with machine-readable (meta)data |  |  | String |  | Optional | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Microstructure |  | Grain size | If polycrystal |  | µm | Numeric | 50 | Mandatory | Rows 51 to 64 apply to the considered example of single crystal Ni-Based Superalloys. This row (51) is an exemplary entry if row 31 = Polycrystal and Microstructure feature (Row 44) = Grain
Metadata | Material history and condition | Microstructure |  | Grain size - Determination method |  |  |  | Drop-down list | Line Intercept / Circular Intercept / Other (Specify) | Mandatory | An exemplary entry if row 31 = Polycrystal and Microstructure feature (Row 44) = Grain
Metadata | Material history and condition | Microstructure |  | Grain size - measuring region | E.g., in the bulk |  |  | String | In the bulk | Mandatory | An exemplary entry if row 31 = Polycrystal and Microstructure feature (Row 44) = Grain
Metadata | Material history and condition | Microstructure |  | Grain size distribution | Link to file, preferably with machine-readable (meta)data |  |  | String | Link(s) to image(s) | Optional | An exemplary entry if row 31 = Polycrystal and Microstructure feature (Row 44) = Grain
Metadata | Material history and condition | Microstructure |  | Grain - Additional Information | Include all further relevant characterization results |  |  | String |  | Optional | An exemplary entry if row 31 = Polycrystal and Microstructure feature (Row 44) = Grain
Metadata | Material history and condition | Microstructure |  | Reprecipitated gamma-gamma prime regions | Completely dissolved and re-precipitated gamma-gamma' regions |  | % | String | ≥ 85 % vol. fraction | Optional | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Microstructure |  | Gamma prime particles - average size |  |  | µm | String | 0.5 ± 0.2 | Optional | An exemplary entry for "Precipitate" (Row 44)
Metadata | Material history and condition | Microstructure |  | Gamma prime particles - maximum size |  |  | µm | Numeric | 1.0 | Optional | An exemplary entry for "Precipitate" (Row 44)
Metadata | Material history and condition | Microstructure |  | Dendrite spacings | Please describe the procedure and add a link to file, preferably with machine-readable (meta)data or add a reference to paper and section |  |  | String | Description of procedure and links to images and raw data | Optional | An exemplary entry for "Dendrite" (Row 44)
Metadata | Material history and condition | Microstructure |  | Proof of single crystallinity | Link to file, preferably with machine-readable (meta)data |  |  | String | Link(s) to X-Ray or EBSD Results | Optional | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Microstructure |  | Single crystal orientation | Link to file, preferably with machine-readable (meta)data. Laue Crystal Verification. Must be documented for each test piece. |  |  | String | Link to analysis results in form of table | Mandatory | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Microstructure |  | Single crystal orientation - Determination method |  |  |  | String | EBSD and XRD for each blank | Mandatory | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Microstructure |  | Single crystal orientation - Measuring point |  |  |  | String | EBSD and XRD for each blank | Mandatory | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Microstructure |  | Orientation - Determination accuracy |  |  |  | String |  | Optional | An exemplary entry for "Other (Specify)" (Row 44)
Metadata | Material history and condition | Chemical composition |  | Measured condition |  |  |  | Drop-down list | As-manufactured / Heat-treated / As-tested / Other (Specify) | Mandatory
Metadata | Material history and condition | Chemical composition |  | Measuring position | Specify position(s) for bulk (middle/top/bottom/surface) and/or microstructure feature (matrix, interdendtitic region, phase,...) |  |  | String |  | Mandatory
Metadata | Material history and condition | Chemical composition |  | Chemical composition - nominal | Link to file, preferably with machine-readable (meta)data or add the wt.-% value of for each element |  | wt.-% / at.-% | String | Link to File | Mandatory
Metadata | Material history and condition | Chemical composition |  | Chemical composition - measured | Include precision, if available. Link to file, preferably with machine-readable (meta)data or add the wt.-% value of for each element |  | wt.-% / at.-% | String | Link to File | Mandatory
Metadata | Material history and condition | Chemical composition |  | Measurement method | Provide a description of the used method(s) and details about measurement volume/points |  |  | String | X-ray fluorescence spectroscopy (XRF) was used for all major elements except for Re. Re contents were determined using inductive coupled plasma atomic emission spectroscopy (ICP-AES) | Mandatory
Metadata | Material history and condition | NDT Results |  | Crack or defect inspection - Method | E.g., Penetrant certification/Radiographic certification/XCT/X-Ray film. Link to file, preferably with machine-readable (meta)data |  |  | String | Link to analysis protocol(s) | Optional
Metadata | Material history and condition | NDT Results |  | Crack or defect inspection - Result |  |  |  | String | Link to analysis protocol(s) | Optional
Metadata | Material history and condition | Mechanical tests results |  | Proof strength - room temperature | 0.2 % Proof strength at room temperature |  | MPa | Numeric | 300 | Optional
Metadata | Material history and condition | Mechanical tests results |  | Proof strength - creep test temperature | 0.2 % Proof strength at creep test temperature |  | MPa | Numeric | 150 | Optional
Metadata | Material history and condition | Mechanical tests results |  | Hardness |  |  |  | String | 400 HV1 | Optional
Metadata | Test piece |  |  | Test piece ID |  |  |  | String | Vh5205-C-99 | Optional
Metadata | Test piece |  |  | Workshop order ID |  |  |  | String |  | Optional
Metadata | Test piece |  |  | Test piece history | Link to file, preferably with machine-readable (meta)data. The file(s) can include, e.g., data or documentation from previous experiments. |  |  | String | Link to data or documentation from previous experiments | Optional
Metadata | Test piece |  |  | Test piece type I |  |  |  | Drop-down list | Miniaturized specimen / Specimen according to standard | Mandatory
Metadata | Test piece |  |  | Test piece type II |  |  |  | Drop-down list | Round cross section / Rectangular section | Mandatory
Metadata | Test piece |  |  | Test piece type III |  |  |  | Drop-down list | Smooth test piece / Notched test piece / Combined test piece / Other type (specify) | Mandatory
Metadata | Test piece |  |  | Test piece technical drawing | Link to file, preferably with machine-readable (meta)data |  |  | String | Link to technical drawing file | Mandatory
Metadata | Test piece |  |  | Test piece origin and orientation | Describe the exact location / positioning of the test piece within the as-tested material. E.g.: Do rolling direction and the longitudinal axis coincide?. Is the test piece located on top or bottom. Do coordinate systems of as-tested material and test piece coincide?. Add a link to a file, e.g., a technical drawing showing this, preferably with machine-readable (meta)data. |  |  | String | Centers of gravity of blanks and test pieces coincide, link to file | Mandatory
Metadata | Test piece |  |  | Test piece orientation in test machine | Describe the exact orientation of the test piece within the test machine. E.g., is the longitudinal axis of the test piece exactly parallel to the loading axis of the test machine?. Add a link to a file providing evidence, preferably with machine-readable (meta)data. This can be, e.g., a technical drawing. |  |  | String |  | Mandatory
Metadata | Test piece |  |  | Additional information test piece | Add the information or a link to file, preferably with machine-readable (meta)data, e.g., Roughness |  |  | String | Roughness Ra = 1 | Optional
Metadata | Measuring and test equipment | Test machine |  | Test machine ID |  |  |  | String |  | Optional
Metadata | Measuring and test equipment | Test machine |  | Test machine type |  |  |  | Drop-down list | Lever arm / Electromechanical drive / Other (Specify) | Optional
Metadata | Measuring and test equipment | Test machine |  | Minimum applicable force |  |  | kN | Numeric | 1 | Optional
Metadata | Measuring and test equipment | Test machine |  | Maximum applicable force |  |  | kN | Numeric | 20 | Optional
Metadata | Measuring and test equipment | Test machine |  | Test frame and specimen alignment | Verification of Test Frame and Specimen Alignment according to ASTM E1012? |  |  | Drop-down list | Yes / No | Mandatory | Entry is mandatory If ASTM E139 (row 7): "Yes" is mandatory. If ISO 204 (row 7): "Yes" is not mandatory. However ISO 204 states that "The test force shall be applied along the test axis in such a manner to minimize bending and torsion of the test piece."
Metadata | Measuring and test equipment | Test machine |  | Test frame and specimen alignment - description | Please provide a description on the procedure followed for the Verification of Test Frame and Specimen Alignment if different from ASTM E1012 |  |  | String |  | Mandatory | Mandatory if verification was not performed according to ASTM E1012 (row 89)
Metadata | Measuring and test equipment | Test machine |  | Calibration class | Calibration class of test frame and specimen alignment, e.g., 5 / 5 starting from 3 kN (according to ASTM E1012) |  |  | String |  | Mandatory | Mandatory if verification according to ASTM E1012 was performed (row 89)
Metadata | Measuring and test equipment | Test machine | Heating system | Furnace type |  |  |  | Drop-down list | Split Tube Furnace with Two-zones / Split Tube Furnace with Three-zones / Other (specify) | Mandatory
Metadata | Measuring and test equipment | Test machine | Test piece holder system | Fixing technique |  |  |  | Drop-down list | Threaded / Button-head / Other (specify) | Mandatory
Metadata | Measuring and test equipment | Test machine | Loading system | Verification of loading system | Was the loading system calibrated/verified? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory if (i) lever-arm machine (row 86) or (ii) uncalibrated load sensor were used (row 108)
Metadata | Measuring and test equipment | Test machine | Loading system | Calibration certificate | Link to file, preferably with machine-readable (meta)data. Please specify which device/feature/part of the test machine was calibrated. |  |  | String | Link to digital certificate that was valid for the test or the test campaigne | Optional
Metadata | Measuring and test equipment | Test machine | Loading system | Calibration date |  |  |  | Date |  | Mandatory
Metadata | Measuring and test equipment | Test machine | Loading system | Calibration validity time period |  |  |  | String |  | Mandatory
Metadata | Measuring and test equipment | Test machine | Loading system | Calibration standard |  |  |  | Drop-down list | DIN EN ISO 7500-2 / Others (specify) | Mandatory
Metadata | Measuring and test equipment | Test machine | Loading system | Calibration class |  |  |  | Numeric | 1 | Mandatory | ISO 204 (row 7): at least class 1 (ISO 7500-2)
Metadata | Measuring and test equipment | Test machine | Loading system | Calibration range |  |  | kN | String | 1-20 | Mandatory
Metadata | Measuring and test equipment | Test machine | Loading system | Use of calibrated weights | Were calibrated weights used to apply the test force? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory in case of a lever-arm machine (row 86) ASTM E139 and ISO 204/ISO 7500-2 set different requirements
Metadata | Measuring and test equipment | Test machine | Loading system | Description of the loading system | Please describe the loading system and related verification in the case that calibrated weights were not used. |  |  | Drop-down list | Yes / No | Mandatory | Mandatory if row 101 = "No"
Metadata | Measuring and test equipment | Test machine | Data acquisition | Data acquisition unit - Model information |  |  |  | String | National Instruments NI SXCI-1600 | Optional
Metadata | Measuring and test equipment | Test machine | Data acquisition | Data acquisition unit - ID |  |  |  | String | 19.08 | Optional
Metadata | Measuring and test equipment | Test machine | Data acquisition | Data acquisition software and version |  |  |  | String | Lab's own development (v22.08.22) | Mandatory
Metadata | Measuring and test equipment | Test machine | Data acquisition | Data acquisition - description |  |  |  | String | Data acquisition was performed every 5 min before load application, every 0.1 s during load application, and every 5 s to max. 10 min after load application, depending on the test duration and when a selected displacement interval is exceeded. | Mandatory
Metadata | Measuring and test equipment | Test machine | Data acquisition | Data acquisition time check | Was the time checked during data acquisition? |  |  | Drop-down list | Yes / No | Optional
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Load sensor during loading | Was a load sensor used during loading? |  |  | Drop-down list | Yes / No | Mandatory | Entry is mandatory. "Yes" is not mandatory.
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Load sensor calibration | Was the load sensor calibrated? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory if row 108 = Yes
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Calibration certificate | Link to file, preferably with machine-readable (meta)data. |  |  | String | Link to digital certificate that was valid for the test or the test campaigne | Optional
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Calibration date |  |  |  | Date |  | Mandatory
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Calibration validity time period |  |  |  | String |  | Mandatory
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Calibration standard |  |  |  | String |  | Mandatory
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Calibration class |  |  |  | Numeric | 1 | Mandatory
Metadata | Measuring and test equipment | Load-measuring system | Load sensor | Calibration Range |  |  | kN | String |  | Mandatory
Metadata | Measuring and test equipment | Load-measuring system | Data acquisition | Force recording | Was the force recorded continuously or periodically (e.g. during loading)? |  |  | String | Phase-wise (during loading) | Mandatory
Metadata | Measuring and test equipment | Laboratory conditions |  | Room temperature | Was the room temperature recorded and checked? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory
Metadata | Measuring and test equipment | Laboratory conditions |  | Room humidity | Was the room humidity recorded and checked? |  |  | Drop-down list | Yes / No | Optional
Metadata | Measuring and test equipment | Temperature-measuring system |  | Temperature signal | Which temperature signal was used for temperature control? |  |  | Drop-down list | On test piece / Via furnace / Other (specify) | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system |  | Metrological traceability | Yes, if temperature sensor and data acquisition are calibrated. |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Sensor type |  |  |  | Drop-down list | Thermocouple / Thermocamera | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Sensor ID |  |  |  | String | 52-PM102-0331 | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Thermocouple type |  |  |  | Drop-down list | B / R / S / K / Other (Specify)… | Mandatory | Rows 123 to 136 are an example for Thermocouple sensor type
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Wire gauge |  |  | mm | Numeric | 0.5 | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Layout | E.g., wire with 2-hole ceramic beads |  |  | String | Wire with 2-hole ceramic beads | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Calibration status | Is/are the thermocouples calibrated? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Calibration method |  |  |  | Drop-down list | Comparison method / Fix-point method | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Calibration standard |  |  |  | Drop-down list | ASTM E220 / EURAMET cg-08 / Guideline DKD-R 5-3 / Other (Specify) | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Calibration certificate | Link to file, preferably with machine-readable (meta)data. |  |  | String |  | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Calibration date |  |  |  | Date |  | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Calibration validity time period |  |  |  | String |  | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Temperature deviation | Measurement deviation detected during calibration |  | °C | String | ± 1.5 | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Calibration Range |  |  | °C | String | 300-1200 | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Contact method |  |  |  | Drop-down list | Welded / Attached / Pressed / Glued / Other (specify) | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Number of thermocouples |  |  |  | Drop-down list | 1 / 2 / 3 | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | Thermocouple location | Location with respect to gauge section |  |  | Drop-down list | Inside / Outside | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Temperature sensor | [7 empty] | Additional fields for non-contacting thermal measurement methods will be provided in upcoming versions
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Calibration status | Is/are the data acquisition unit calibrated? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Reference junction |  |  |  | Drop-down list | Internal / External | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Calibration certificate | Link to file, preferably with machine-readable (meta)data. |  |  | String | Link to digital certificate that was valid for the test or the test campaigne | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Calibration date |  |  |  | Date |  | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Calibration validity time period |  |  |  | String |  | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Calibration method |  |  |  | String | Comparison method | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Calibration standard |  |  |  | Drop-down list | EURAMET cg-11 / Other (specify) | Mandatory
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Temperature deviation | Measurement deviation detected during calibration |  | °C | String | ± 1.5 | Optional
Metadata | Measuring and test equipment | Temperature-measuring system | Data acquisition | Calibration Range |  |  | °C | String | 100-1200 | Mandatory
Metadata | Measuring and test equipment | Extensometer system |  | Displacement measuring method | Type of strain measuring device |  |  | Drop-down list | Contacting method / Non-contacting method | Mandatory | Rows 150 to 167 are an example for a contacting extensometer
Metadata | Measuring and test equipment | Extensometer system |  | Sensor type - Contacting method |  |  |  | Drop-down list | LVDT with extension legs / Clip-on extensometer / Analog dial gauges / Other (Specify) | Mandatory | Applies if Row 147 = Contacting method
Metadata | Measuring and test equipment | Extensometer system |  | Sensor type - Non-contacting method |  |  |  | Drop-down list | Laserextensometer / Digital image correlation technique / Other (Specify) | Mandatory | Applies if Row 147 = Non-contacting method
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Measurement set-up | Measurement one-sided or two-sided? |  |  | Drop-down list | One-sided / Two-sided | Mandatory | ASTM E139 requires two-sided strain measurement
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Extension averaging | Was there an averaging of the extension values? (two-sided extensometer) |  |  | Drop-down list | Yes / No | Mandatory | Entry is mandatory in case of two-sided extensometer systems (row 150). "Yes" is not mandatory.
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Measurement direction |  |  |  | Drop-down list | Axial action / Diametrical action | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Mounting type |  |  |  | Drop-down list | At test piece in parallel length / At collars / Other (Specify) | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Extensometer model information |  |  |  | String | Answers could be, e.g., MTS 632.53F-31 or HBM WA20 | Optional
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Extensometer ID | The ID used for identification in the laboratory |  |  | String | 19.05 | Optional
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Extensometer leg material | Material of upper/lower legs, e.g., in LVDT systems |  |  | Drop-down list | Ceramic legs / Metallic legs | Optional
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Measuring amplifier - Model information |  |  |  | String |  | Optional
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Extension range - Upper limit |  |  | % / mm | Numeric | Answers could be, e.g., 20 mm, or 5 mm, or -20 mm | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Extension range - Lower limit |  |  | % / mm | Numeric | Answers could be, e.g., -10 mm, -2,5 mm, or 0 mm | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Nominal gauge length | If applicable |  | mm | Numeric | 25 | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Is the extensometer incl. the data acquisition calibrated? | Is the extensometer incl. the data acquisition calibrated? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Calibration certificate | Link to file, preferably with machine-readable (meta)data. |  |  | String | Link to digital certificate that was valid for the test or the test campaigne | Optional
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Calibration date |  |  |  | Date |  | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Calibration validity time period |  |  |  | String |  | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Calibration class |  |  |  | String | Answers could be, e.g., 1, or B2 | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Calibration Range |  |  | % (mm) | String | 30 (7.5) | Mandatory
Metadata | Measuring and test equipment | Extension values | Contacting extensometer | Calibration standard |  |  |  | String | Answers could be, e.g., DIN EN ISO 9513 or ASTM E83 | Mandatory
Metadata | Measuring and test equipment | Extension values | Non-contacting extensometer | [7 empty] | Additional fields for non-contacting strain measurement methods will be provided in upcoming versions
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Measuring equipment |  |  |  | Drop-down list | Caliper gauge / Measuring microscope / Micrometer screw gauge / Other (specify) | Mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Measuring equipment's usage |  |  |  | String | Determination of elongation characteristic values | Mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Resolution |  |  | mm | String | 0.01 | Mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Type |  |  |  | Drop-down list | Digital / Analog / Other (specify) | Optional
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Equipment model information |  |  |  | String | Mitutoyo CD-15WW | Optional
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Equipment ID | The ID used for identification in the laboratory |  |  | String | 47.06 | Optional
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Calibration status | Is the measuring equipment calibrated? |  |  | Drop-down list | Yes / No | Mandatory | "Yes" is mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Calibration certificate | Link to file, preferably with machine-readable (meta)data. |  |  | String | Link to digital certificate that was valid for the test or the test campaigne | Optional
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Calibration date |  |  |  | Date |  | Mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Calibration validity time period |  |  |  | String | 12 months | Mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Calibration result |  |  |  | Drop-down list | Meets calibration requirements / Does not meet calibration requirements | Mandatory | "Meets calibration requirements" is mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Calibration range |  |  | mm | Numeric | 150 | Mandatory
Metadata | Measuring and test equipment | Elongation values and cross-sectional dimensions |  | Calibration standard |  |  |  | String | Answers could be, e.g., VDI, VDE, or DGQ 2618 Blatt 9.1 | Mandatory
Metadata | Data processing procedures |  |  | Primary data series | Primary data is data that is directly acquired by sensors or measuring instruments during or after a test. Please add a list of the measured quantities and their corresponsing units. |  |  | String |  | Mandatory
Metadata | Data processing procedures |  |  | Processed data series | Processed data is obtained as a result of using procedures (equations, algorithms, methods, unit conversions, averaging, smoothing) to transform primary data. Please describe the transformed quantities, their corresponsing units, and the applied procedures. |  |  | String |  | Mandatory
Metadata | Data processing procedures |  |  | Data analysis procedures | Description of the data processing and analysis procedures used to obtain specific test results, e.g. percentage elastic extension, ee |  |  | String |  | Mandatory
Metadata | Data processing procedures |  |  | Workflow usage | Were automated (user-independent) analysis workflows used? |  |  | Drop-down list | Yes / No | Mandatory | Entry is mandatory. "Yes" is not mandatory.
Metadata | Data processing procedures |  |  | Software | If applicable, please list the used software/workflow, including product and version |  |  | Drop-down list | Not applicable / Free text field Other (Specify) | Mandatory
Metadata | Data processing procedures |  |  | Related publications | If applicable, please list publications related to the data analysis procedure/software used |  |  | String, Array |  | Optional
Primary data | Test result | Values recorded at test start |  | Minimum test piece diameter at room temperature |  | D | mm | Numeric | 6003 | Mandatory | If round cross section (test piece, row 79)
Primary data | Test result | Values recorded at test start |  | Original gauge length |  | Lo | mm | Numeric | 30 | Mandatory
Primary data | Test result | Values recorded at test start |  | Parallel length |  | Lc | mm | Numeric |  | Mandatory
Primary data | Test result | Values recorded at test start |  | Extensometer gauge length |  | Le | mm | Numeric | 22.78 | Mandatory
Primary data | Test result | Values recorded at test start |  | Determination of reference length | Case 1: Reference length for calculation of percentage extensions/elongations if the original gauge length Lo and/or extensometer Le are inside the parallel length, Lc. In this case Lr = Le (percentage extensions) or Lr = Lo (percentage elongations). Case 2: Reference length for calculation of percentage extensions/elongations if the original gauge length Lo and/or extensometer Le are outside the parallel length, Lc. In this case, Lr, for calculation of percentage is corrected to consider the strain contributions of the shoulders/ridges. |  |  | Drop-down list | Case 1 / Case 2 | Mandatory
Primary data | Test result | Values recorded at test start |  | Reference length to calculate percentage elongations |  | Lr | mm | Numeric | 30 | Mandatory
Primary data | Test result | Values recorded at test start |  | Reference length to calculate percentage extensions |  | Lr | mm | Numeric | 22.78 | Mandatory
Primary data | Test result | Values recorded at test start |  | k-Value | Lr / √So |  | mm | Numeric | 5.65 | Mandatory | If ISO 204 (row 7)
Primary data | Test result | Values recorded at test start |  | Ratio reference length to diameter | Lr / D |  | mm | Numeric | Typical answers could be, e.g., 4 or 5 | Mandatory | If round cross section (test piece, row 79)
Primary data | Test result | Values recorded during test run |  | Elapsed time from end of loading | Link to file, preferably with machine-readable (meta)data. | t | s / h | String |  | Mandatory
Primary data | Test result | Values recorded during test run |  | Force |  |  | kN | Numeric |  | Mandatory | Only required during loading if the test was performed under constant force (row 10)
Primary data | Test result | Values recorded during test run |  | Extension | Link to file, preferably with machine-readable (meta)data. | ΔLet | mm | String |  | Optional
Primary data | Test result | Values recorded during test run |  | Elongation | Link to file, preferably with machine-readable (meta)data. | ΔLot | mm | String |  | Optional
Primary data | Test result | Values recorded during test run |  | Heating time |  |  | h | Numeric | 2.97 | Mandatory
Primary data | Test result | Values recorded during test run |  | Soak time | Soak time before the test |  | h | Numeric | 1.25 | Mandatory
Primary data | Test result | Values recorded during test run |  | Test duration |  | t | h | Numeric | 166.3 | Mandatory
Primary data | Test result | Values recorded after end of test |  | Creep rupture time |  | tu | h | Numeric | 166 | Mandatory | Mandatory if test performed until test piece break (row 13)
Primary data | Test result | Values recorded after end of test |  | Fracture position |  |  |  | Drop-down list | Centered / Within the gauge length / Within the gauge length and the parallel length / At the edge of the upper extensometer gauge length / At the edge of the lower extensometer gauge length / At the edge of the upper parallel length / At the edge of the lower parallel length / Outside the extensometer gauge length / Outside the extensometer the extensometer gauge length and the parallel length | Mandatory | Mandatory if test performed until test piece break (row 13)
Primary data | Test result | Values recorded after end of test |  | Final gauge length after fracture |  | Lu | mm | Numeric |  | Optional
Secondary data | Test result | Values recorded during test run |  | Corrected measured temperature | Link to file, preferably with machine-readable (meta)data. | Tc | ℃ | String | Link to file | Mandatory
Secondary data | Test result | Values recorded during test run |  | Loading rate |  |  | MPa/s | Numeric | 11.33 | Optional
Secondary data | Test result | Values recorded during test run |  | Unloading rate |  |  | MPa/s | Numeric |  | Optional
Secondary data | Test result | Values recorded during test run |  | Heating speed |  |  | ℃/min | Numeric | 5.4 | Optional
Secondary data | Test result | Values recorded during test run |  | Cooling speed |  |  | ℃/min | Numeric |  | Optional
Secondary data | Test result | Values recorded during test run |  | Percentage extension | Link to file, preferably with machine-readable (meta)data. Data series corresponds to the percentage plastic extension from end of loading. If the percentage initial plastic extension is not zero, these data series coresponds to the percentage creep extension from end of loading. | e | % | String |  | Mandatory | Extension values mandatory if continuous monitoring of extension (first option in row 12)
Secondary data | Test result | Elongation values |  | Percentage permanent elongation |  | Aper | % | Numeric |  | Mandatory | Mandatory only if the test was not run until test piece break (row 13)
Secondary data | Test result | Elongation values |  | Percentage elongation after creep fracture |  | Au | % | Numeric |  | Mandatory | With or without extension mearurement
Secondary data | Test result | Elongation values |  | Percentage reduction of area after creep fracture |  | Zu | % | Numeric |  | Mandatory | With or without extension mearurement
Secondary data | Test result | Extension values |  | Percentage total extension |  | et | % | Numeric | 28.0 | Mandatory
Secondary data | Test result | Extension values |  | Percentage initial total extension |  | eti | % | Numeric |  | Mandatory
Secondary data | Test result | Extension values |  | Percentage elastic extension |  | ee | % | Numeric |  | Mandatory
Secondary data | Test result | Extension values |  | Percentage initial plastic extension |  | ei | % | Numeric | 0 | Mandatory
Secondary data | Test result | Extension values |  | Percentage plastic extension |  | ep | % | Numeric | 27.7 | Mandatory
Secondary data | Test result | Extension values |  | Percentage creep extension |  | ef | % | Numeric | 27.7 | Mandatory